
Product overview
The handheld dimmable attenuator is a precise, durable, portable instrument specifically designed to install, operate, and maintain fiber optic networks. Features a small form factor, a backlit display with selectable switches, an ultra-wide light attenuation range, precise attenuation accuracy, and user-calibrating functions and a universal interface design, while the attenuation value is displayed via the LCD screen.
Product features
It uses a rechargeable lithium battery for up to 48 hours of continuous operating time
The attenuation value is adjustable
The LCD display corresponds to the attenuation value
The unique FC/SC/ST common interface (shown in Figures 1 and 2) eliminates the need for complex conversions
Optional auto-shutdown function
Optional backlit display with on/off
Product annotation diagram
[image: image1.jpg]


Figure1，2

Technical specifications
	Gauge model
	TYPE A
	TYPE B

	Attenuation range（DB）
	0-30
	0～60

	Uncertainty
	±5%

	Operating temperature(℃)
	-10～+60

	Storage temperature(℃)
	-25～+70

	Automatic shutdown time(min)
	10

	The battery lasts for continuous operating time
	Not less than 48 hours

	Dimensions(mm)
	190×100×48

	Power supply
	Rechargeable lithium battery

	Weight(g)
	400

	Insertion loss
	<0.8dB

	Maximum input optical power
	500mW


Function description
1、 LCD Display
   LCD displays light attenuation impairment, displayed in the form of dB; The set wavelengths are 1310 nm, 1550 nm.
2、 ON/OFF Button
Press the button to power on the system when the system is powered off.
Press the button briefly (less than 3 seconds) in the power-on state to turn on or off the auto-shutdown function. If the "Auto Shutdown Flag" is displayed, the auto shutdown function is turned on, and the auto shutdown will be automatically after 10 minutes; If it is not displayed, the auto-shutdown function is turned off. 

Press and hold (greater than or equal to 3 seconds) to turn off the light attenuator while the system is on.
3、 UP Button
     Each time this button is pressed, the attenuation value increases by one step value.
4、 LIGHT Button
   Pressing the button allows you to turn the backlight of the LCD on and off. In working mode, press LIGHT to turn the backlight on/off. When the backlight is turned on, a small sun sign appears on the top left of the screen, indicating that the backlight is turned on.
5、 λButton
   λ wavelength selection button, press this button, you can choose different wavelengths, there are 1310 nm, 1550 nm six wavelengths for selection, this value will also be displayed on the LCD。

6、 DOWN Button
   Each time this button is pressed, the attenuation value decreases by one step value.
7、 STEP Button
   Press this button to attenuate the attenuation step in a loop between 0.1DB, 1DB, 10DB.
8、 Serial port debugging function
   Connect the serial port connection end with the PC serial port terminal, and the instrument attenuation value can be set by the PC (as shown in the figure on the right).
9、 [image: image2.jpg]Auto shutdown
sign

Backlit sign
Attenuation
impairment

‘Wavelength sign

‘UP Button
Step button

Abutton

'ON/OFF Button
LIGHT Button

DOWN Button.



Rechargeable function
     During normal operation, the power indicator sign is not displayed; When the battery is low, the battery indicator flashes every second, and will automatically shut down after ten flashes; At this time, the charger directly connected to the instrument can be charged (as shown in the figure on the right), which can save the trouble of replacing the battery and related costs.
Instructions for use
   As shown in Figures 1, 2, the input light at the IN end, the out output light, and the desired optical attenuation value is set by the UP and DOWN buttons.
Standard configuration
	Configuration description

	Dimmable attenuator host
Rechargeable lithium battery
Product warranty card
Oxford Soft Pack
charger
Data Cable (Option)
Certificate


Maintenance and upkeep
1. Keep the sensor end face clean frequently, so that it is grease-free and pollution-free, do not use unclean, non-standard adapter connectors, and do not insert the end face with poor polishing surface, otherwise it will damage the sensor end face and greatly reduce the performance of the entire system.
2. Stick to one adapter whenever possible.
3. Once the dimmable attenuator is not in use, the dust cap is immediately covered to protect the end surface from long-term exposure to dust and measurement errors caused by long-term exposure to dust in the air.
4. Carefully insert and remove the optical adapter connector to avoid scratches on the port.
5. Clean the sensor surface regularly. When cleaning the sensor surface, use a special cleaning swab to gently wipe in the circumferential direction.
Common failures
	Fault performance
	Possible causes
	Workaround

	The LCD display is weak
	Insufficient power supply
	charge

	No display at power-on
	Insufficient power supply/other
	Power on/charge again

	The LCD display data remains unchanged or changes slightly
	The optical adapter connector is faulty or dirty
The turbidity/display is locked
	Detect optical adapter connector connections Correct;

 Clean the sensor face


Quality assurance
We don’t endorse user repair of smart handheld dimmable attenuators.
1. The warranty period of the instrument is 18 months from the date of shipment. Our company will be all its products material and process commitment, warranty period of eighteen months from the date of shipment is valid. When the purchased product is found to have quality problems during this period, our company will make corresponding repairs or replacements. However, in any case, our company's liability will not exceed the purchase price of the product.
1. If there is a problem in the use of the instrument, according to the common fault prompt solution can not be solved, the user is not allowed to open the chassis without authorization, please contact our company's marketing department or local agents.
2. For quality failures caused by production defects, our company is responsible for free repair or replacement of products.
This guarantee applies only to normal use of the meter and is free from damage or improper use.
The warranty of the product does not cover problems/failures caused by:
1) Unauthorized repair or modification of the instrument
2) Improper use, negligent use, accident, etc
All my products have a warranty registration card. Please fill in and send it back to our company together with a copy of the invoice, so that when we need to maintain, innovate, calibrate, etc. of your instrument in the future, we will have a root cause record.

